## Quant Summary
Quantitative | Analysis | - | Summary | Report
Sample | ID: | Cal Blank
Sample | Date/Time: | Thursday, April 20, 2021 21:38:30
Sample | Description:
Solution | Type: | Blank
Blank | File: | C:\Path\Tp\Your\Data\Cal Blank.001
Number | of | Replicates: | 3
Peak | Processing | Mode: | Average
Signal | Profile | Processing | Mode: | Average
Dual | Detector | Mode: | Dual
Dead | Time | (ns): | 55
Sample | File: | C:\Path\To\Your\Data\Sample\04-20-21 CAL & SAMPLES.sam
Method | File: | C:\Path\To\Your\Data\Method\heavymetals incl is.mth
Dataset | File: | C:\Path\To\Your\Data\04202021\Cal Blank.001
Tuning | File: | C:\Path\To\Your\Data\Tuning\default.tun
Optimization | File: | C:\Path\To\Your\Data\Optimize\default.dac
Calibration | File: | C:\Path\To\Your\Data\System\04202021.cal
Calibration | Type: | External Calibration
Summary
Intensities
| Analyte | Mass | Meas. Intens. Mean | Meas. Intens. RSD | Blank Intensity | Blank Intens. RSD
¦> | Ge | 74 | 47281 | 0.748
¦- | As | 75 | 65 | 4.448
¦- | Hg | 202 | 52 | 9.174
¦ | Pb | 208 | 247 | 4.597
¦> | Bi | 209 | 419010 | 1.227
¦- | Cd | 114 | 226 | 13.932
¦> | In | 115 | 491446 | 1.26
Concentration | Results
| Analyte | Mass | Net Intens. Mean | Conc. Mean | Conc. SD | Conc. RSD | Sample Unit
¦> | Ge | 74 | ppb
¦- | As | 75 | ppb
¦- | Hg | 202 | ppb
¦ | Pb | 208 | ppb
¦> | Bi | 209 | ppb
¦- | Cd | 114 | ppb
¦> | In | 115 | ppb
Quantitative | Analysis | - | Summary | Report
Sample | ID: | Cal 1 0.25ppb 4 mix
Sample | Date/Time: | Thursday, June 03, 2021 21:41:36
Sample | Description:
Solution | Type: | Standard
Blank | File: | C:\Path\To\Your\Data\DataSet\04202021\Cal Blank.001
Number | of | Replicates: | 3
Peak | Processing | Mode: | Average
Signal | Profile | Processing | Mode: | Average
Dual | Detector | Mode: | Dual
Dead | Time | (ns): | 55
Sample | File: | C:\Path\To\Your\Data\Sample\04-20-21 CAL & SAMPLES.sam
Method | File: | C:\Path\To\Your\Data\Method\heavymetals incl is.mth
Dataset | File: | C:\Path\To\Your\Data\04202021\Cal 1 0.25ppb 4 mix.002
Tuning | File: | C:\Path\To\Your\Data\Tuning\default.tun
Optimization | File: | C:\Path\To\Your\Data\Optimize\default.dac
Calibration | File: | C:\Path\To\Your\Data\System\04202021.cal
Calibration | Type: | External Calibration
Summary
Intensities
| Analyte | Mass | Meas. Intens. Mean | Meas. Intens. RSD | Blank Intensity | Blank Intens. RSD
¦> | Ge | 74 | 46954 | 0.645 | 47281.288 | 0.748
¦- | As | 75 | 285 | 6.904 | 65.334 | 4.448
¦- | Hg | 202 | 962 | 2.861 | 51.556 | 9.174
¦ | Pb | 208 | 4588 | 1.662 | 247.113 | 4.597
¦> | Bi | 209 | 426286 | 0.771 | 419009.917 | 1.227
¦- | Cd | 114 | 3326 | 1.434 | 225.657 | 13.932
¦> | In | 115 | 484214 | 1.136 | 491446.115 | 1.26
Concentration | Results
| Analyte | Mass | Net Intens. Mean | Conc. Mean | Conc. SD | Conc. RSD | Sample Unit
¦> | Ge | 74 | 46954.386 | ppb
¦- | As | 75 | 0.005 | 0.222 | 0.02 | 9 | ppb
¦- | Hg | 202 | 0.002 | 0.209 | 0.01 | 3.5 | ppb
¦ | Pb | 208 | 0.01 | 0.208 | 0 | 1 | ppb
¦> | Bi | 209 | 426285.993 | ppb
¦- | Cd | 114 | 0.006 | 0.229 | 0 | 1 | ppb
¦> | In | 115 | 484214.232 | ppb
Quantitative | Analysis | - | Summary | Report
Sample | ID: | Cal 2 0.5ppb 4 mix
Sample | Date/Time: | Thursday, June 03, 2021 21:44:39
Sample | Description:
Solution | Type: | Standard
Blank | File: | C:\Path\To\Your\Data\04202021\Cal Blank.001
Number | of | Replicates: | 3
Peak | Processing | Mode: | Average
Signal | Profile | Processing | Mode: | Average
Dual | Detector | Mode: | Dual
Dead | Time | (ns): | 55
Sample | File: | C:\Path\To\Your\Data\Sample\04-20-21 CAL & SAMPLES.sam
Method | File: | C:\Path\To\Your\Data\Method\heavymetals incl is.mth
Dataset | File: | C:\Path\To\Your\Data\04202021\Cal 2 0.5ppb 4 mix.003
Tuning | File: | C:\Path\To\Your\Data\Tuning\default.tun
Optimization | File: | C:\Path\To\Your\Data\Optimize\default.dac
Calibration | File: | C:\Path\To\Your\Data\System\04202021.cal
Calibration | Type: | External Calibration
Summary
Intensities
| Analyte | Mass | Meas. Intens. Mean | Meas. Intens. RSD | Blank Intensity | Blank Intens. RSD
¦> | Ge | 74 | 46891 | 0.312 | 47281.288 | 0.748
¦- | As | 75 | 490 | 2.264 | 65.334 | 4.448
¦- | Hg | 202 | 1866 | 2.699 | 51.556 | 9.174
¦ | Pb | 208 | 9357 | 1.069 | 247.113 | 4.597
¦> | Bi | 209 | 428395 | 0.71 | 419009.917 | 1.227
¦- | Cd | 114 | 6404 | 1.592 | 225.657 | 13.932
¦> | In | 115 | 486655 | 1.724 | 491446.115 | 1.26
Concentration | Results
| Analyte | Mass | Net Intens. Mean | Conc. Mean | Conc. SD | Conc. RSD | Sample Unit
¦> | Ge | 74 | 46890.761 | ppb
¦- | As | 75 | 0.009 | 0.428 | 0.01 | 2.5 | ppb
¦- | Hg | 202 | 0.004 | 0.415 | 0.01 | 2.1 | ppb
¦ | Pb | 208 | 0.021 | 0.435 | 0 | 0.4 | ppb
¦> | Bi | 209 | 428394.559 | ppb
¦- | Cd | 114 | 0.013 | 0.454 | 0.01 | 2.1 | ppb
¦> | In | 115 | 486655.063 | ppb
Quantitative | Analysis | - | Summary | Report
Sample | ID: | Cal 3 2.5ppb 4 mix
Sample | Date/Time: | Thursday, June 03, 2021 21:47:43
Sample | Description:
Solution | Type: | Standard
Blank | File: | C:\Path\To\Your\Data\04202021\Cal Blank.001
Number | of | Replicates: | 3
Peak | Processing | Mode: | Average
Signal | Profile | Processing | Mode: | Average
Dual | Detector | Mode: | Dual
Dead | Time | (ns): | 55
Sample | File: | C:\Path\To\Your\Data\Sample\04-20-21 CAL & SAMPLES.sam
Method | File: | C:\Path\To\Your\Data\Method\heavymetals incl is.mth
Dataset | File: | C:\Path\To\Your\Data\04202021\Cal 3 2.5ppb 4 mix.004
Tuning | File: | C:\Path\To\Your\Data\Tuning\default.tun
Optimization | File: | C:\Path\To\Your\Data\Optimize\default.dac
Calibration | File: | C:\Path\To\Your\Data\System\04202021.cal
Calibration | Type: | External Calibration
Summary
Intensities
| Analyte | Mass | Meas. Intens. Mean | Meas. Intens. RSD | Blank Intensity | Blank Intens. RSD
¦> | Ge | 74 | 46776 | 0.952 | 47281.288 | 0.748
¦- | As | 75 | 2330 | 2.339 | 65.334 | 4.448
¦- | Hg | 202 | 9919 | 1.545 | 51.556 | 9.174
¦ | Pb | 208 | 47767 | 1.413 | 247.113 | 4.597
¦> | Bi | 209 | 427832 | 1.117 | 419009.917 | 1.227
¦- | Cd | 114 | 31689 | 0.89 | 225.657 | 13.932
¦> | In | 115 | 485197 | 0.85 | 491446.115 | 1.26
Concentration | Results
| Analyte | Mass | Net Intens. Mean | Conc. Mean | Conc. SD | Conc. RSD | Sample Unit
¦> | Ge | 74 | 46776.245 | ppb
¦- | As | 75 | 0.048 | 2.285 | 0.07 | 3.1 | ppb
¦- | Hg | 202 | 0.023 | 2.26 | 0.03 | 1.4 | ppb
¦ | Pb | 208 | 0.111 | 2.273 | 0.06 | 2.4 | ppb
¦> | Bi | 209 | 427832.111 | ppb
¦- | Cd | 114 | 0.065 | 2.317 | 0.03 | 1.1 | ppb
¦> | In | 115 | 485197.344 | ppb
Quantitative | Analysis | - | Summary | Report
Sample | ID: | Cal 4 5ppb 4 mix
Sample | Date/Time: | Thursday, June 03, 2021 21:50:47
Sample | Description:
Solution | Type: | Standard
Blank | File: | C:\Path\To\Your\Data\04202021\Cal Blank.001
Number | of | Replicates: | 3
Peak | Processing | Mode: | Average
Signal | Profile | Processing | Mode: | Average
Dual | Detector | Mode: | Dual
Dead | Time | (ns): | 55
Sample | File: | C:\Path\To\Your\Data\Sample\04-20-21 CAL & SAMPLES.sam
Method | File: | C:\Path\To\Your\Data\Method\heavymetals incl is.mth
Dataset | File: | C:\Path\To\Your\Data\04202021\Cal 4 5ppb 4 mix.005
Tuning | File: | C:\Path\To\Your\Data\Tuning\default.tun
Optimization | File: | C:\Path\To\Your\Data\Optimize\default.dac
Calibration | File: | C:\Path\To\Your\Data\System\04202021.cal
Calibration | Type: | External Calibration
Summary
Intensities
| Analyte | Mass | Meas. Intens. Mean | Meas. Intens. RSD | Blank Intensity | Blank Intens. RSD
¦> | Ge | 74 | 46606 | 0.12 | 47281.288 | 0.748
¦- | As | 75 | 4701 | 1.861 | 65.334 | 4.448
¦- | Hg | 202 | 20639 | 0.428 | 51.556 | 9.174
¦ | Pb | 208 | 98834 | 0.919 | 247.113 | 4.597
¦> | Bi | 209 | 429093 | 1.11 | 419009.917 | 1.227
¦- | Cd | 114 | 64810 | 0.335 | 225.657 | 13.932
¦> | In | 115 | 487959 | 0.814 | 491446.115 | 1.26
Concentration | Results
| Analyte | Mass | Net Intens. Mean | Conc. Mean | Conc. SD | Conc. RSD | Sample Unit
¦> | Ge | 74 | 46605.858 | ppb
¦- | As | 75 | 0.099 | 4.693 | 0.09 | 1.9 | ppb
¦- | Hg | 202 | 0.048 | 4.702 | 0.03 | 0.7 | ppb
¦ | Pb | 208 | 0.23 | 4.701 | 0.04 | 0.8 | ppb
¦> | Bi | 209 | 429092.821 | ppb
¦- | Cd | 114 | 0.132 | 4.73 | 0.04 | 0.8 | ppb
¦> | In | 115 | 487958.591 | ppb
Quantitative | Analysis | - | Summary | Report
Sample | ID: | Cal 5 10ppb 4 mix
Sample | Date/Time: | Thursday, June 03, 2021 21:53:50
Sample | Description:
Solution | Type: | Standard
Blank | File: | C:\Path\To\Your\Data\04202021\Cal Blank.001
Number | of | Replicates: | 3
Peak | Processing | Mode: | Average
Signal | Profile | Processing | Mode: | Average
Dual | Detector | Mode: | Dual
Dead | Time | (ns): | 55
Sample | File: | C:\Path\To\Your\Data\Sample\04-20-21 CAL & SAMPLES.sam
Method | File: | C:\Path\To\Your\Data\Method\heavymetals incl is.mth
Dataset | File: | C:\Path\To\Your\Data\04202021\Cal 5 10ppb 4 mix.006
Tuning | File: | C:\Path\To\Your\Data\Tuning\default.tun
Optimization | File: | C:\Path\To\Your\Data\Optimize\default.dac
Calibration | File: | C:\Path\To\Your\Data\System\04202021.cal
Calibration | Type: | External Calibration
Summary
Intensities
| Analyte | Mass | Meas. Intens. Mean | Meas. Intens. RSD | Blank Intensity | Blank Intens. RSD
¦> | Ge | 74 | 46321 | 1.225 | 47281.288 | 0.748
¦- | As | 75 | 9788 | 0.445 | 65.334 | 4.448
¦- | Hg | 202 | 43555 | 1.152 | 51.556 | 9.174
¦ | Pb | 208 | 204685 | 0.491 | 247.113 | 4.597
¦> | Bi | 209 | 426637 | 1.649 | 419009.917 | 1.227
¦- | Cd | 114 | 134078 | 0.474 | 225.657 | 13.932
¦> | In | 115 | 483298 | 1.807 | 491446.115 | 1.26
Concentration | Results
| Analyte | Mass | Net Intens. Mean | Conc. Mean | Conc. SD | Conc. RSD | Sample Unit
¦> | Ge | 74 | 46321.142 | ppb
¦- | As | 75 | 0.21 | 9.905 | 0.09 | 0.9 | ppb
¦- | Hg | 202 | 0.102 | 9.994 | 0.09 | 0.9 | ppb
¦ | Pb | 208 | 0.479 | 9.805 | 0.11 | 1.1 | ppb
¦> | Bi | 209 | 426636.533 | ppb
¦- | Cd | 114 | 0.277 | 9.898 | 0.15 | 1.6 | ppb
¦> | In | 115 | 483298.434 | ppb
Quantitative | Analysis | - | Summary | Report
Sample | ID: | Cal 6 15ppb 4 mix
Sample | Date/Time: | Thursday, June 03, 2021 21:56:55
Sample | Description:
Solution | Type: | Standard
Blank | File: | C:\Path\To\Your\Data\04202021\Cal Blank.001
Number | of | Replicates: | 3
Peak | Processing | Mode: | Average
Signal | Profile | Processing | Mode: | Average
Dual | Detector | Mode: | Dual
Dead | Time | (ns): | 55
Sample | File: | C:\Path\To\Your\Data\Sample\04-20-21 CAL & SAMPLES.sam
Method | File: | C:\Path\To\Your\Data\Method\heavymetals incl is.mth
Dataset | File: | C:\Path\To\Your\Data\04202021\Cal 6 15ppb 4 mix.007
Tuning | File: | C:\Path\To\Your\Data\Tuning\default.tun
Optimization | File: | C:\Path\To\Your\Data\Optimize\default.dac
Calibration | File: | C:\Path\To\Your\Data\System\04202021.cal
Calibration | Type: | External Calibration
Summary
Intensities
| Analyte | Mass | Meas. Intens. Mean | Meas. Intens. RSD | Blank Intensity | Blank Intens. RSD
¦> | Ge | 74 | 46487 | 0.562 | 47281.288 | 0.748
¦- | As | 75 | 15045 | 0.436 | 65.334 | 4.448
¦- | Hg | 202 | 66878 | 1.437 | 51.556 | 9.174
¦ | Pb | 208 | 322954 | 1.206 | 247.113 | 4.597
¦> | Bi | 209 | 432390 | 0.109 | 419009.917 | 1.227
¦- | Cd | 114 | 205256 | 0.644 | 225.657 | 13.932
¦> | In | 115 | 482296 | 0.443 | 491446.115 | 1.26
Concentration | Results
| Analyte | Mass | Net Intens. Mean | Conc. Mean | Conc. SD | Conc. RSD | Sample Unit
¦> | Ge | 74 | 46487.303 | ppb
¦- | As | 75 | 0.322 | 15.204 | 0.09 | 0.6 | ppb
¦- | Hg | 202 | 0.155 | 15.147 | 0.22 | 1.5 | ppb
¦ | Pb | 208 | 0.746 | 15.27 | 0.19 | 1.3 | ppb
¦> | Bi | 209 | 432389.542 | ppb
¦- | Cd | 114 | 0.425 | 15.191 | 0.16 | 1.1 | ppb
¦> | In | 115 | 482296.06 | ppb
Quantitative | Analysis | - | Summary | Report
Sample | ID: | 2.5ppb CCV 4 MIX
Sample | Date/Time: | Thursday, June 03, 2021 22:00:01
Sample | Description: | 2.5CCV
Solution | Type: | Sample
Blank | File: | C:\Path\To\Your\Data\04202021\Cal Blank.001
Number | of | Replicates: | 3
Peak | Processing | Mode: | Average
Signal | Profile | Processing | Mode: | Average
Dual | Detector | Mode: | Dual
Dead | Time | (ns): | 55
Sample | File: | C:\Path\To\Your\Data\Sample\04-20-21 CAL & SAMPLES.sam
Method | File: | C:\Path\To\Your\Data\Method\heavymetals incl is.mth
Dataset | File: | C:\Path\To\Your\Data\04202021\2.5ppb CCV 4 MIX.008
Tuning | File: | C:\Path\To\Your\Data\Tuning\default.tun
Optimization | File: | C:\Path\To\Your\Data\Optimize\default.dac
Calibration | File: | C:\Path\To\Your\Data\System\04202021.cal
Calibration | Type: | External Calibration
Summary
Intensities
| Analyte | Mass | Meas. Intens. Mean | Meas. Intens. RSD | Blank Intensity | Blank Intens. RSD
¦> | Ge | 74 | 46593 | 0.742 | 47281.288 | 0.748
¦- | As | 75 | 2341 | 2.042 | 65.334 | 4.448
¦- | Hg | 202 | 10558 | 0.813 | 51.556 | 9.174
¦ | Pb | 208 | 48899 | 1.657 | 247.113 | 4.597
¦> | Bi | 209 | 431323 | 0.262 | 419009.917 | 1.227
¦- | Cd | 114 | 31841 | 1.03 | 225.657 | 13.932
¦> | In | 115 | 483025 | 0.703 | 491446.115 | 1.26
Concentration | Results
| Analyte | Mass | Net Intens. Mean | Conc. Mean | Conc. SD | Conc. RSD | Sample Unit | CCV Recovery (%)
¦> | Ge | 74 | 46592.945 | ppb
¦- | As | 75 | 0.049 | 2.306 | 0.06 | 2.6 | ppb | 92.24
¦- | Hg | 202 | 0.024 | 2.387 | 0.01 | 0.6 | ppb | 95.48
¦ | Pb | 208 | 0.113 | 2.307 | 0.03 | 1.4 | ppb | 92.28
¦> | Bi | 209 | 431323.112 | ppb
¦- | Cd | 114 | 0.065 | 2.339 | 0.03 | 1.3 | ppb | 93.56
¦> | In | 115 | 483024.751 | ppb
Quantitative | Analysis | - | Summary | Report
Sample | ID: | Digestion Blank 06022021
Sample | Date/Time: | Thursday, June 03, 2021 22:03:06
Sample | Description:
Solution | Type: | Sample
Blank | File: | C:\Path\To\Your\Data\04202021\Cal Blank.001
Number | of | Replicates: | 3
Peak | Processing | Mode: | Average
Signal | Profile | Processing | Mode: | Average
Dual | Detector | Mode: | Dual
Dead | Time | (ns): | 55
Sample | File: | C:\Path\To\Your\Data\Sample\04-20-21 CAL & SAMPLES.sam
Method | File: | C:\Path\To\Your\Data\Method\heavymetals incl is.mth
Dataset | File: | C:\Path\To\Your\Data\04202021\Digestion Blank 06022021.009
Tuning | File: | C:\Path\To\Your\Data\Tuning\default.tun
Optimization | File: | C:\Path\To\Your\Data\Optimize\default.dac
Calibration | File: | C:\Path\To\Your\Data\System\04202021.cal
Calibration | Type: | External Calibration
Summary
Intensities
| Analyte | Mass | Meas. Intens. Mean | Meas. Intens. RSD | Blank Intensity | Blank Intens. RSD
¦> | Ge | 74 | 48088 | 0.138 | 47281.288 | 0.748
¦- | As | 75 | 168 | 1.657 | 65.334 | 4.448
¦- | Hg | 202 | 769 | 29.9 | 51.556 | 9.174
¦ | Pb | 208 | 266 | 23.656 | 247.113 | 4.597
¦> | Bi | 209 | 429211 | 1.19 | 419009.917 | 1.227
¦- | Cd | 114 | 277 | 13.902 | 225.657 | 13.932
¦> | In | 115 | 502259 | 1.893 | 491446.115 | 1.26
Concentration | Results
| Analyte | Mass | Net Intens. Mean | Conc. Mean | Conc. SD | Conc. RSD | Sample Unit
¦> | Ge | 74 | 48087.872 | ppb
¦- | As | 75 | 0.002 | 0.099 | 0 | 2.7 | ppb
¦- | Hg | 202 | 0.002 | 0.164 | 0.05 | 33.2 | ppb
¦ | Pb | 208 | 0 | 0.001 | 0 | 499.5 | ppb
¦> | Bi | 209 | 429210.525 | ppb
¦- | Cd | 114 | 0 | 0.003 | 0 | 71.4 | ppb
¦> | In | 115 | 502258.714 | ppb
Quantitative | Analysis | - | Summary | Report
Sample | ID: | Digestion Blank 06022021 + spike
Sample | Date/Time: | Thursday, June 03, 2021 22:18:24
Sample | Description: | 100 uL of 1 ppm mix - 2 ppb
Solution | Type: | Sample
Blank | File: | C:\Path\To\Your\Data\04202021\Cal Blank.001
Number | of | Replicates: | 3
Peak | Processing | Mode: | Average
Signal | Profile | Processing | Mode: | Average
Dual | Detector | Mode: | Dual
Dead | Time | (ns): | 55
Sample | File: | C:\Path\To\Your\Data\Sample\04-20-21 CAL & SAMPLES.sam
Method | File: | C:\Path\To\Your\Data\Method\heavymetals incl is.mth
Dataset | File: | C:\Path\To\Your\Data\04202021\Digestion Blank 06022021 + spike.014
Tuning | File: | C:\Path\To\Your\Data\Tuning\default.tun
Optimization | File: | C:\Path\To\Your\Data\Optimize\default.dac
Calibration | File: | C:\Path\To\Your\Data\System\04202021.cal
Calibration | Type: | External Calibration
Summary
Intensities
| Analyte | Mass | Meas. Intens. Mean | Meas. Intens. RSD | Blank Intensity | Blank Intens. RSD
¦> | Ge | 74 | 48237 | 0.847 | 47281.288 | 0.748
¦- | As | 75 | 2339 | 1.482 | 65.334 | 4.448
¦- | Hg | 202 | 8814 | 0.296 | 51.556 | 9.174
¦ | Pb | 208 | 43409 | 0.31 | 247.113 | 4.597
¦> | Bi | 209 | 428117 | 0.386 | 419009.917 | 1.227
¦- | Cd | 114 | 28844 | 1.145 | 225.657 | 13.932
¦> | In | 115 | 514373 | 1.292 | 491446.115 | 1.26
Concentration | Results
| Analyte | Mass | Net Intens. Mean | Conc. Mean | Conc. SD | Conc. RSD | Sample Unit
¦> | Ge | 74 | 48237.135 | ppb
¦- | As | 75 | 0.047 | 2.223 | 0.05 | 2.3 | ppb
¦- | Hg | 202 | 0.02 | 2.006 | 0.01 | 0.4 | ppb
¦ | Pb | 208 | 0.101 | 2.063 | 0 | 0.1 | ppb
¦> | Bi | 209 | 428116.579 | ppb
¦- | Cd | 114 | 0.056 | 1.987 | 0.04 | 1.8 | ppb
¦> | In | 115 | 514372.937 | ppb
Quantitative | Analysis | - | Summary | Report
Sample | ID: | QC Std 1
Sample | Date/Time: | Thursday, June 03, 2021 22:30:46
Sample | Description:
Solution | Type: | QC Std
Blank | File: | C:\Path\To\Your\Data\04202021\Cal Blank.001
Number | of | Replicates: | 3
Peak | Processing | Mode: | Average
Signal | Profile | Processing | Mode: | Average
Dual | Detector | Mode: | Dual
Dead | Time | (ns): | 55
Sample | File: | C:\Path\To\Your\Data\Sample\04-20-21 CAL & SAMPLES.sam
Method | File: | C:\Path\To\Your\Data\Method\heavymetals incl is.mth
Dataset | File: | C:\Path\To\Your\Data\04202021\QC Std 1.018
Tuning | File: | C:\Path\To\Your\Data\Tuning\default.tun
Optimization | File: | C:\Path\To\Your\Data\Optimize\default.dac
Calibration | File: | C:\Path\To\Your\Data\System\04202021.cal
Calibration | Type: | External Calibration
Summary
Intensities
| Analyte | Mass | Meas. Intens. Mean | Meas. Intens. RSD | Blank Intensity | Blank Intens. RSD
¦> | Ge | 74 | 48121 | 1.857 | 47281.288 | 0.748
¦- | As | 75 | 536 | 1.458 | 65.334 | 4.448
¦- | Hg | 202 | 1927 | 1.339 | 51.556 | 9.174
¦ | Pb | 208 | 9511 | 1.407 | 247.113 | 4.597
¦> | Bi | 209 | 443198 | 0.877 | 419009.917 | 1.227
¦- | Cd | 114 | 6549 | 2.499 | 225.657 | 13.932
¦> | In | 115 | 505286 | 1.784 | 491446.115 | 1.26
Concentration | Results
| Analyte | Mass | Net Intens. Mean | Conc. Mean | Conc. SD | Conc. RSD | Sample Unit | CCV Recovery (%)
¦> | Ge | 74 | 48121.061 | ppb
¦- | As | 75 | 0.01 | 0.46 | 0.02 | 3.3 | ppb | 92
¦- | Hg | 202 | 0.004 | 0.414 | 0 | 0.7 | ppb | 82.8
¦ | Pb | 208 | 0.021 | 0.427 | 0 | 0.6 | ppb | 85.4
¦> | Bi | 209 | 443197.659 | ppb
¦- | Cd | 114 | 0.013 | 0.447 | 0.02 | 4.4 | ppb | 89.4
¦> | In | 115 | 505286.111 | ppb
Quantitative | Analysis | - | Summary | Report
Sample | ID: | QC Std 1
Sample | Date/Time: | Thursday, June 03, 2021 22:34:12
Sample | Description:
Solution | Type: | QC Std
Blank | File: | C:\Path\To\Your\Data\04202021\Cal Blank.001
Number | of | Replicates: | 3
Peak | Processing | Mode: | Average
Signal | Profile | Processing | Mode: | Average
Dual | Detector | Mode: | Dual
Dead | Time | (ns): | 55
Sample | File: | C:\Path\To\Your\Data\Sample\04-20-21 CAL & SAMPLES.sam
Method | File: | C:\Path\To\Your\Data\Method\heavymetals incl is.mth
Dataset | File: | C:\Path\To\Your\Data\04202021\QC Std 1.019
Tuning | File: | C:\Path\To\Your\Data\Tuning\default.tun
Optimization | File: | C:\Path\To\Your\Data\Optimize\default.dac
Calibration | File: | C:\Path\To\Your\Data\System\04202021.cal
Calibration | Type: | External Calibration
Summary
Intensities
| Analyte | Mass | Meas. Intens. Mean | Meas. Intens. RSD | Blank Intensity | Blank Intens. RSD
¦> | Ge | 74 | 45840 | 1.305 | 47281.288 | 0.748
¦- | As | 75 | 526 | 3.269 | 65.334 | 4.448
¦- | Hg | 202 | 1912 | 0.634 | 51.556 | 9.174
¦ | Pb | 208 | 9500 | 0.355 | 247.113 | 4.597
¦> | Bi | 209 | 436780 | 1.584 | 419009.917 | 1.227
¦- | Cd | 114 | 6347 | 1.133 | 225.657 | 13.932
¦> | In | 115 | 491683 | 1.27 | 491446.115 | 1.26
Concentration | Results
| Analyte | Mass | Net Intens. Mean | Conc. Mean | Conc. SD | Conc. RSD | Sample Unit | CCV Recovery (%)
¦> | Ge | 74 | 45840.217 | ppb
¦- | As | 75 | 0.01 | 0.477 | 0.02 | 3.9 | ppb | 95.4
¦- | Hg | 202 | 0.004 | 0.417 | 0 | 1.1 | ppb | 83.4
¦ | Pb | 208 | 0.021 | 0.433 | 0.01 | 1.8 | ppb | 86.6
¦> | Bi | 209 | 436779.516 | ppb
¦- | Cd | 114 | 0.012 | 0.445 | 0 | 0.2 | ppb | 89
¦> | In | 115 | 491683.486 | ppb
Quantitative | Analysis | - | Summary | Report
Sample | ID: | QC Std 3
Sample | Date/Time: | Thursday, June 03, 2021 22:39:46
Sample | Description:
Solution | Type: | QC Std
Blank | File: | C:\Path\To\Your\Data\04202021\Cal Blank.001
Number | of | Replicates: | 3
Peak | Processing | Mode: | Average
Signal | Profile | Processing | Mode: | Average
Dual | Detector | Mode: | Dual
Dead | Time | (ns): | 55
Sample | File: | C:\Path\To\Your\Data\Sample\04-20-21 CAL & SAMPLES.sam
Method | File: | C:\Path\To\Your\Data\Method\heavymetals incl is.mth
Dataset | File: | C:\Path\To\Your\Data\04202021\QC Std 3.020
Tuning | File: | C:\Path\To\Your\Data\Tuning\default.tun
Optimization | File: | C:\Path\To\Your\Data\Optimize\default.dac
Calibration | File: | C:\Path\To\Your\Data\System\04202021.cal
Calibration | Type: | External Calibration
Summary
Intensities
| Analyte | Mass | Meas. Intens. Mean | Meas. Intens. RSD | Blank Intensity | Blank Intens. RSD
¦> | Ge | 74 | 44749 | 0.601 | 47281.288 | 0.748
¦- | As | 75 | 2248 | 2.118 | 65.334 | 4.448
¦- | Hg | 202 | 10308 | 1.147 | 51.556 | 9.174
¦ | Pb | 208 | 49928 | 0.807 | 247.113 | 4.597
¦> | Bi | 209 | 451423 | 1.169 | 419009.917 | 1.227
¦- | Cd | 114 | 30833 | 0.76 | 225.657 | 13.932
¦> | In | 115 | 483127 | 1.537 | 491446.115 | 1.26
Concentration | Results
| Analyte | Mass | Net Intens. Mean | Conc. Mean | Conc. SD | Conc. RSD | Sample Unit | CCV Recovery (%)
¦> | Ge | 74 | 44748.547 | ppb
¦- | As | 75 | 0.049 | 2.305 | 0.06 | 2.7 | ppb | 92.2
¦- | Hg | 202 | 0.023 | 2.226 | 0.05 | 2.3 | ppb | 89.04
¦ | Pb | 208 | 0.11 | 2.251 | 0.01 | 0.4 | ppb | 90.04
¦> | Bi | 209 | 451423.488 | ppb
¦- | Cd | 114 | 0.063 | 2.264 | 0.02 | 0.8 | ppb | 90.56
¦> | In | 115 | 483127.169 | ppb
Quantitative | Analysis | - | Summary | Report
Sample | ID: | Digestion Blank 06022021 2
Sample | Date/Time: | Thursday, June 03, 2021 22:50:22
Sample | Description:
Solution | Type: | Sample
Blank | File: | C:\Path\To\Your\Data\04202021\Cal Blank.001
Number | of | Replicates: | 3
Peak | Processing | Mode: | Average
Signal | Profile | Processing | Mode: | Average
Dual | Detector | Mode: | Dual
Dead | Time | (ns): | 55
Sample | File: | C:\Path\To\Your\Data\Sample\04-20-21 CAL & SAMPLES.sam
Method | File: | C:\Path\To\Your\Data\Method\heavymetals incl is.mth
Dataset | File: | C:\Path\To\Your\Data\04202021\Digestion Blank 06022021 2.022
Tuning | File: | C:\Path\To\Your\Data\Tuning\default.tun
Optimization | File: | C:\Path\To\Your\Data\Optimize\default.dac
Calibration | File: | C:\Path\To\Your\Data\System\04202021.cal
Calibration | Type: | External Calibration
Summary
Intensities
| Analyte | Mass | Meas. Intens. Mean | Meas. Intens. RSD | Blank Intensity | Blank Intens. RSD
¦> | Ge | 74 | 49412 | 2.244 | 47281.288 | 0.748
¦- | As | 75 | 224 | 3.694 | 65.334 | 4.448
¦- | Hg | 202 | 64 | 8.136 | 51.556 | 9.174
¦ | Pb | 208 | 243 | 3.593 | 247.113 | 4.597
¦> | Bi | 209 | 454033 | 0.47 | 419009.917 | 1.227
¦- | Cd | 114 | 196 | 4.421 | 225.657 | 13.932
¦> | In | 115 | 536278 | 1.367 | 491446.115 | 1.26
Concentration | Results
| Analyte | Mass | Net Intens. Mean | Conc. Mean | Conc. SD | Conc. RSD | Sample Unit
¦> | Ge | 74 | 49412.447 | ppb
¦- | As | 75 | 0.003 | 0.149 | 0.01 | 8.4 | ppb
¦- | Hg | 202 | 0 | 0.002 | 0 | 67.3 | ppb
¦ | Pb | 208 | 0 | -0.001 | 0 | 31.5 | ppb
¦> | Bi | 209 | 454032.597 | ppb
¦- | Cd | 114 | 0 | -0.003 | 0 | 14.1 | ppb
¦> | In | 115 | 536277.618 | ppb
Quantitative | Analysis | - | Summary | Report
Sample | ID: | KLS-1223
Sample | Date/Time: | Thursday, June 03, 2021 23:02:56
Sample | Description:
Solution | Type: | Sample
Blank | File: | C:\Path\To\Your\Data\04202021\Cal Blank.001
Number | of | Replicates: | 3
Peak | Processing | Mode: | Average
Signal | Profile | Processing | Mode: | Average
Dual | Detector | Mode: | Dual
Dead | Time | (ns): | 55
Sample | File: | C:\Path\To\Your\Data\Sample\04-20-21 CAL & SAMPLES.sam
Method | File: | C:\Path\To\Your\Data\Method\heavymetals incl is.mth
Dataset | File: | C:\Path\To\Your\Data\04202021\KLS-1223.026
Tuning | File: | C:\Path\To\Your\Data\Tuning\default.tun
Optimization | File: | C:\Path\To\Your\Data\Optimize\default.dac
Calibration | File: | C:\Path\To\Your\Data\System\04202021.cal
Calibration | Type: | External Calibration
Summary
Intensities
| Analyte | Mass | Meas. Intens. Mean | Meas. Intens. RSD | Blank Intensity | Blank Intens. RSD
¦> | Ge | 74 | 55226 | 0.641 | 47281.288 | 0.748
¦- | As | 75 | 1146 | 2.738 | 65.334 | 4.448
¦- | Hg | 202 | 96 | 8.526 | 51.556 | 9.174
¦ | Pb | 208 | 3641 | 1.196 | 247.113 | 4.597
¦> | Bi | 209 | 431845 | 0.895 | 419009.917 | 1.227
¦- | Cd | 114 | 2254 | 1.4 | 225.657 | 13.932
¦> | In | 115 | 527665 | 0.98 | 491446.115 | 1.26
Concentration | Results
| Analyte | Mass | Net Intens. Mean | Conc. Mean | Conc. SD | Conc. RSD | Sample Unit | Sample Dilution Ratio | Final Sample Concentration (ppm) | Conc Calc For Spike Recovery (ppm)
¦> | Ge | 74 | 55225.876 | ppb
¦- | As | 75 | 0.019 | 0.914 | 0.03 | 2.9 | ppb | 102.5 | 0.093685 | 0.093685
¦- | Hg | 202 | 0 | 0.01 | 0 | 18.7 | ppb | 102.5 | <LOQ | 0.001025
¦ | Pb | 208 | 0.008 | 0.16 | 0 | 0.6 | ppb | 102.5 | <LOQ | 0.0164
¦> | Bi | 209 | 431844.593 | ppb
¦- | Cd | 114 | 0.004 | 0.136 | 0 | 0.6 | ppb | 102.5 | <LOQ | 0.01394
¦> | In | 115 | 527664.501 | ppb
Quantitative | Analysis | - | Summary | Report
Sample | ID: | KLS-1223 + SPIKE
Sample | Date/Time: | Thursday, June 03, 2021 23:06:01
Sample | Description: | 100 uL of 1 ppm mix - 2 ppb
Solution | Type: | Sample
Blank | File: | C:\Path\To\Your\Data\04202021\Cal Blank.001
Number | of | Replicates: | 3
Peak | Processing | Mode: | Average
Signal | Profile | Processing | Mode: | Average
Dual | Detector | Mode: | Dual
Dead | Time | (ns): | 55
Sample | File: | C:\Path\To\Your\Data\Sample\04-20-21 CAL & SAMPLES.sam
Method | File: | C:\Path\To\Your\Data\Method\heavymetals incl is.mth
Dataset | File: | C:\Path\To\Your\Data\04202021\KLS-1223 + SPIKE.027
Tuning | File: | C:\Path\To\Your\Data\Tuning\default.tun
Optimization | File: | C:\Path\To\Your\Data\Optimize\default.dac
Calibration | File: | C:\Path\To\Your\Data\System\04202021.cal
Calibration | Type: | External Calibration
Summary
Intensities
| Analyte | Mass | Meas. Intens. Mean | Meas. Intens. RSD | Blank Intensity | Blank Intens. RSD
¦> | Ge | 74 | 56229 | 0.534 | 47281.288 | 0.748
¦- | As | 75 | 4686 | 2.026 | 65.334 | 4.448
¦- | Hg | 202 | 8377 | 0.143 | 51.556 | 9.174
¦ | Pb | 208 | 46078 | 0.359 | 247.113 | 4.597
¦> | Bi | 209 | 428071 | 0.205 | 419009.917 | 1.227
¦- | Cd | 114 | 31939 | 0.396 | 225.657 | 13.932
¦> | In | 115 | 532486 | 1.114 | 491446.115 | 1.26
Concentration | Results
| Analyte | Mass | Net Intens. Mean | Conc. Mean | Conc. SD | Conc. RSD | Sample Unit | Sample Dilution Ratio | Final Sample Concentration (ppm) | Final Spike Recovery %
¦> | Ge | 74 | 56228.549 | ppb
¦- | As | 75 | 0.082 | 3.867 | 0.1 | 2.5 | ppb | 102.5 | 0.39637 | 147.65
¦- | Hg | 202 | 0.019 | 1.906 | 0.01 | 0.3 | ppb | 102.5 | 0.19536 | 94.8
¦ | Pb | 208 | 0.107 | 2.19 | 0 | 0.2 | ppb | 102.5 | 0.22448 | 101.5
¦> | Bi | 209 | 428070.878 | ppb
¦- | Cd | 114 | 0.06 | 2.127 | 0.03 | 1.3 | ppb | 102.5 | 0.21802 | 99.55
¦> | In | 115 | 532485.501 | ppb
Quantitative | Analysis | - | Summary | Report
Sample | ID: | Digestion Blank 04202021
Sample | Date/Time: | Thursday, June 03, 2021 23:09:07
Sample | Description: | Digestion Blank 04202021
Solution | Type: | Sample
Blank | File: | C:\Path\To\Your\Data\04202021\Cal Blank.001
Number | of | Replicates: | 3
Peak | Processing | Mode: | Average
Signal | Profile | Processing | Mode: | Average
Dual | Detector | Mode: | Dual
Dead | Time | (ns): | 55
Sample | File: | C:\Path\To\Your\Data\Sample\04-20-21 CAL & SAMPLES.sam
Method | File: | C:\Path\To\Your\Data\Method\heavymetals incl is.mth
Dataset | File: | C:\Path\To\Your\Data\04202021\Digestion Blank 04202021.028
Tuning | File: | C:\Path\To\Your\Data\Tuning\default.tun
Optimization | File: | C:\Path\To\Your\Data\Optimize\default.dac
Calibration | File: | C:\Path\To\Your\Data\System\04202021.cal
Calibration | Type: | External Calibration
Summary
Intensities
| Analyte | Mass | Meas. Intens. Mean | Meas. Intens. RSD | Blank Intensity | Blank Intens. RSD
¦> | Ge | 74 | 51679 | 1.778 | 47281.288 | 0.748
¦- | As | 75 | 254 | 4.719 | 65.334 | 4.448
¦- | Hg | 202 | 105 | 13.062 | 51.556 | 9.174
¦ | Pb | 208 | 289 | 21.258 | 247.113 | 4.597
¦> | Bi | 209 | 471222 | 1.134 | 419009.917 | 1.227
¦- | Cd | 114 | 252 | 16.881 | 225.657 | 13.932
¦> | In | 115 | 535785 | 0.753 | 491446.115 | 1.26
Concentration | Results
| Analyte | Mass | Net Intens. Mean | Conc. Mean | Conc. SD | Conc. RSD | Sample Unit
¦> | Ge | 74 | 51678.687 | ppb
¦- | As | 75 | 0.004 | 0.167 | 0.01 | 4.7 | ppb
¦- | Hg | 202 | 0 | 0.01 | 0 | 27.7 | ppb
¦ | Pb | 208 | 0 | 0 | 0 | 532.2 | ppb
¦> | Bi | 209 | 471222.333 | ppb
¦- | Cd | 114 | 0 | 0 | 0 | 676.6 | ppb
¦> | In | 115 | 535784.829 | ppb
Quantitative | Analysis | - | Summary | Report
Sample | ID: | KLS-1224
Sample | Date/Time: | Thursday, June 03, 2021 23:12:13
Sample | Description: | KLS-1224
Solution | Type: | Sample
Blank | File: | C:\Path\To\Your\Data\04202021\Cal Blank.001
Number | of | Replicates: | 3
Peak | Processing | Mode: | Average
Signal | Profile | Processing | Mode: | Average
Dual | Detector | Mode: | Dual
Dead | Time | (ns): | 55
Sample | File: | C:\Path\To\Your\Data\Sample\04-20-21 CAL & SAMPLES.sam
Method | File: | C:\Path\To\Your\Data\Method\heavymetals incl is.mth
Dataset | File: | C:\Path\To\Your\Data\04202021\KLS-1224.029
Tuning | File: | C:\Path\To\Your\Data\Tuning\default.tun
Optimization | File: | C:\Path\To\Your\Data\Optimize\default.dac
Calibration | File: | C:\Path\To\Your\Data\System\04202021.cal
Calibration | Type: | External Calibration
Summary
Intensities
| Analyte | Mass | Meas. Intens. Mean | Meas. Intens. RSD | Blank Intensity | Blank Intens. RSD
¦> | Ge | 74 | 59367 | 0.816 | 47281.288 | 0.748
¦- | As | 75 | 214 | 2.655 | 65.334 | 4.448
¦- | Hg | 202 | 80 | 11.644 | 51.556 | 9.174
¦ | Pb | 208 | 3639 | 0.801 | 247.113 | 4.597
¦> | Bi | 209 | 499074 | 1.479 | 419009.917 | 1.227
¦- | Cd | 114 | 208 | 8.376 | 225.657 | 13.932
¦> | In | 115 | 636929 | 0.887 | 491446.115 | 1.26
Concentration | Results
| Analyte | Mass | Net Intens. Mean | Conc. Mean | Conc. SD | Conc. RSD | Sample Unit | Sample Dilution Ratio | Final Sample Concentration (ppm) | Conc Calc For Spike Recovery (ppm)
¦> | Ge | 74 | 59367.342 | ppb
¦- | As | 75 | 0.002 | 0.105 | 0.01 | 5.1 | ppb | 387.814 | <LOQ | 0.04072
¦- | Hg | 202 | 0 | 0.004 | 0 | 44.4 | ppb | 387.814 | <LOQ | 0.0015513
¦ | Pb | 208 | 0.007 | 0.137 | 0 | 2.2 | ppb | 387.814 | <LOQ | 0.053131
¦> | Bi | 209 | 499073.857 | ppb
¦- | Cd | 114 | 0 | -0.005 | 0 | 20.1 | ppb | 387.814 | <LOQ | -0.0019391
¦> | In | 115 | 636928.579 | ppb
Quantitative | Analysis | - | Summary | Report
Sample | ID: | KLS-1224 + SPIKE
Sample | Date/Time: | Thursday, June 03, 2021 23:15:25
Sample | Description: | 100 uL of 1 ppm spike - 2ppb
Solution | Type: | Sample
Blank | File: | C:\Path\To\Your\Data\04202021\Cal Blank.001
Number | of | Replicates: | 3
Peak | Processing | Mode: | Average
Signal | Profile | Processing | Mode: | Average
Dual | Detector | Mode: | Dual
Dead | Time | (ns): | 55
Sample | File: | C:\Path\To\Your\Data\Sample\04-20-21 CAL & SAMPLES.sam
Method | File: | C:\Path\To\Your\Data\Method\heavymetals incl is.mth
Dataset | File: | C:\Path\To\Your\Data\04202021\KLS-1224 + SPIKE.030
Tuning | File: | C:\Path\To\Your\Data\Tuning\default.tun
Optimization | File: | C:\Path\To\Your\Data\Optimize\default.dac
Calibration | File: | C:\Path\To\Your\Data\System\04202021.cal
Calibration | Type: | External Calibration
Summary
Intensities
| Analyte | Mass | Meas. Intens. Mean | Meas. Intens. RSD | Blank Intensity | Blank Intens. RSD
¦> | Ge | 74 | 58651 | 1.215 | 47281.288 | 0.748
¦- | As | 75 | 3890 | 0.5 | 65.334 | 4.448
¦- | Hg | 202 | 9095 | 0.763 | 51.556 | 9.174
¦ | Pb | 208 | 49945 | 0.633 | 247.113 | 4.597
¦> | Bi | 209 | 494092 | 0.694 | 419009.917 | 1.227
¦- | Cd | 114 | 35401 | 1.274 | 225.657 | 13.932
¦> | In | 115 | 626271 | 0.397 | 491446.115 | 1.26
Concentration | Results
| Analyte | Mass | Net Intens. Mean | Conc. Mean | Conc. SD | Conc. RSD | Sample Unit | Sample Dilution Ratio | Final Sample Concentration (ppm) | Final Spike Recovery %
¦> | Ge | 74 | 58651.27 | ppb
¦- | As | 75 | 0.065 | 3.064 | 0.05 | 1.7 | ppb | 282.3193 | 0.86503 | 145.9882
¦- | Hg | 202 | 0.018 | 1.792 | 0.03 | 1.5 | ppb | 282.3193 | 0.50592 | 89.3253
¦ | Pb | 208 | 0.1 | 2.056 | 0.02 | 1.2 | ppb | 282.3193 | 0.58045 | 93.3903
¦> | Bi | 209 | 494091.691 | ppb
¦- | Cd | 114 | 0.056 | 2.003 | 0.02 | 0.9 | ppb | 282.3193 | 0.56549 | 100.4934
¦> | In | 115 | 626270.824 | ppb
Quantitative | Analysis | - | Summary | Report
Sample | ID: | QC Std 3
Sample | Date/Time: | Thursday, June 03, 2021 23:20:48
Sample | Description:
Solution | Type: | QC Std
Blank | File: | C:\Path\To\Your\Data\04202021\Cal Blank.001
Number | of | Replicates: | 3
Peak | Processing | Mode: | Average
Signal | Profile | Processing | Mode: | Average
Dual | Detector | Mode: | Dual
Dead | Time | (ns): | 55
Sample | File: | C:\Path\To\Your\Data\Sample\04-20-21 CAL & SAMPLES.sam
Method | File: | C:\Path\To\Your\Data\Method\heavymetals incl is.mth
Dataset | File: | C:\Path\To\Your\Data\04202021\QC Std 3.031
Tuning | File: | C:\Path\To\Your\Data\Tuning\default.tun
Optimization | File: | C:\Path\To\Your\Data\Optimize\default.dac
Calibration | File: | C:\Path\To\Your\Data\System\04202021.cal
Calibration | Type: | External Calibration
Summary
Intensities
| Analyte | Mass | Meas. Intens. Mean | Meas. Intens. RSD | Blank Intensity | Blank Intens. RSD
¦> | Ge | 74 | 45737 | 1.195 | 47281.288 | 0.748
¦- | As | 75 | 2456 | 0.776 | 65.334 | 4.448
¦- | Hg | 202 | 10683 | 0.606 | 51.556 | 9.174
¦ | Pb | 208 | 51489 | 0.951 | 247.113 | 4.597
¦> | Bi | 209 | 465720 | 1.238 | 419009.917 | 1.227
¦- | Cd | 114 | 31374 | 0.806 | 225.657 | 13.932
¦> | In | 115 | 496076 | 1.455 | 491446.115 | 1.26
Concentration | Results
| Analyte | Mass | Net Intens. Mean | Conc. Mean | Conc. SD | Conc. RSD | Sample Unit | CCV Recovery (%)
¦> | Ge | 74 | 45736.868 | ppb
¦- | As | 75 | 0.052 | 2.469 | 0.03 | 1.4 | ppb | 98.76
¦- | Hg | 202 | 0.023 | 2.236 | 0.02 | 0.7 | ppb | 89.44
¦ | Pb | 208 | 0.11 | 2.25 | 0.04 | 1.9 | ppb | 90
¦> | Bi | 209 | 465719.518 | ppb
¦- | Cd | 114 | 0.063 | 2.244 | 0.04 | 1.8 | ppb | 89.76
¦> | In | 115 | 496076.422 | ppb
Quantitative | Analysis | - | Summary | Report
Sample | ID: | QC Std 3
Sample | Date/Time: | Thursday, June 03, 2021 23:24:18
Sample | Description:
Solution | Type: | QC Std
Blank | File: | C:\Path\To\Your\Data\04202021\Cal Blank.001
Number | of | Replicates: | 3
Peak | Processing | Mode: | Average
Signal | Profile | Processing | Mode: | Average
Dual | Detector | Mode: | Dual
Dead | Time | (ns): | 55
Sample | File: | C:\Path\To\Your\Data\Sample\04-20-21 CAL & SAMPLES.sam
Method | File: | C:\Path\To\Your\Data\Method\heavymetals incl is.mth
Dataset | File: | C:\Path\To\Your\Data\04202021\QC Std 3.032
Tuning | File: | C:\Path\To\Your\Data\Tuning\default.tun
Optimization | File: | C:\Path\To\Your\Data\Optimize\default.dac
Calibration | File: | C:\Path\To\Your\Data\System\04202021.cal
Calibration | Type: | External Calibration
Summary
Intensities
| Analyte | Mass | Meas. Intens. Mean | Meas. Intens. RSD | Blank Intensity | Blank Intens. RSD
¦> | Ge | 74 | 44016 | 1.11 | 47281.288 | 0.748
¦- | As | 75 | 2289 | 2.745 | 65.334 | 4.448
¦- | Hg | 202 | 10658 | 0.526 | 51.556 | 9.174
¦ | Pb | 208 | 51400 | 0.933 | 247.113 | 4.597
¦> | Bi | 209 | 457745 | 0.494 | 419009.917 | 1.227
¦- | Cd | 114 | 30845 | 2.229 | 225.657 | 13.932
¦> | In | 115 | 478762 | 1.818 | 491446.115 | 1.26
Concentration | Results
| Analyte | Mass | Net Intens. Mean | Conc. Mean | Conc. SD | Conc. RSD | Sample Unit | CCV Recovery (%)
¦> | Ge | 74 | 44015.775 | ppb
¦- | As | 75 | 0.051 | 2.388 | 0.04 | 1.8 | ppb | 95.52
¦- | Hg | 202 | 0.023 | 2.27 | 0.02 | 0.7 | ppb | 90.8
¦ | Pb | 208 | 0.112 | 2.285 | 0.02 | 0.9 | ppb | 91.4
¦> | Bi | 209 | 457744.894 | ppb
¦- | Cd | 114 | 0.064 | 2.286 | 0.01 | 0.5 | ppb | 91.44
¦> | In | 115 | 478762.084 | ppb
Quantitative | Analysis | - | Summary | Report
Sample | ID: | KLS-1225
Sample | Date/Time: | Thursday, June 03, 2021 23:27:22
Sample | Description: | KLS-1225
Solution | Type: | Sample
Blank | File: | C:\Path\To\Your\Data\04202021\Cal Blank.001
Number | of | Replicates: | 3
Peak | Processing | Mode: | Average
Signal | Profile | Processing | Mode: | Average
Dual | Detector | Mode: | Dual
Dead | Time | (ns): | 55
Sample | File: | C:\Path\To\Your\Data\Sample\04-20-21 CAL & SAMPLES.sam
Method | File: | C:\Path\To\Your\Data\Method\heavymetals incl is.mth
Dataset | File: | C:\Path\To\Your\Data\04202021\KLS-1225.033
Tuning | File: | C:\Path\To\Your\Data\Tuning\default.tun
Optimization | File: | C:\Path\To\Your\Data\Optimize\default.dac
Calibration | File: | C:\Path\To\Your\Data\System\04202021.cal
Calibration | Type: | External Calibration
Summary
Intensities
| Analyte | Mass | Meas. Intens. Mean | Meas. Intens. RSD | Blank Intensity | Blank Intens. RSD
¦> | Ge | 74 | 55446 | 0.841 | 47281.288 | 0.748
¦- | As | 75 | 234 | 9.101 | 65.334 | 4.448
¦- | Hg | 202 | 141 | 11.044 | 51.556 | 9.174
¦ | Pb | 208 | 3372 | 0.911 | 247.113 | 4.597
¦> | Bi | 209 | 484665 | 1.909 | 419009.917 | 1.227
¦- | Cd | 114 | 391 | 10.081 | 225.657 | 13.932
¦> | In | 115 | 608687 | 1.255 | 491446.115 | 1.26
Concentration | Results
| Analyte | Mass | Net Intens. Mean | Conc. Mean | Conc. SD | Conc. RSD | Sample Unit | Sample Dilution Ratio | Final Sample Concentration (ppm)
¦> | Ge | 74 | 55445.811 | ppb
¦- | As | 75 | 0.003 | 0.134 | 0.02 | 14.5 | ppb | 351 | <LOQ
¦- | Hg | 202 | 0 | 0.016 | 0 | 16.1 | ppb | 351 | <LOQ
¦ | Pb | 208 | 0.006 | 0.13 | 0 | 1.2 | ppb | 351 | <LOQ
¦> | Bi | 209 | 484664.958 | ppb
¦- | Cd | 114 | 0 | 0.007 | 0 | 37.7 | ppb | 351 | <LOQ
¦> | In | 115 | 608687.008 | ppb
Quantitative | Analysis | - | Summary | Report
Sample | ID: | KLS-1226
Sample | Date/Time: | Thursday, June 03, 2021 23:30:26
Sample | Description: | KLS-1226
Solution | Type: | Sample
Blank | File: | C:\Path\To\Your\Data\04202021\Cal Blank.001
Number | of | Replicates: | 3
Peak | Processing | Mode: | Average
Signal | Profile | Processing | Mode: | Average
Dual | Detector | Mode: | Dual
Dead | Time | (ns): | 55
Sample | File: | C:\Path\To\Your\Data\Sample\04-20-21 CAL & SAMPLES.sam
Method | File: | C:\Path\To\Your\Data\Method\heavymetals incl is.mth
Dataset | File: | C:\Path\To\Your\Data\04202021\KLS-1226.034
Tuning | File: | C:\Path\To\Your\Data\Tuning\default.tun
Optimization | File: | C:\Path\To\Your\Data\Optimize\default.dac
Calibration | File: | C:\Path\To\Your\Data\System\04202021.cal
Calibration | Type: | External Calibration
Summary
Intensities
| Analyte | Mass | Meas. Intens. Mean | Meas. Intens. RSD | Blank Intensity | Blank Intens. RSD
¦> | Ge | 74 | 52523 | 0.482 | 47281.288 | 0.748
¦- | As | 75 | 1119 | 1.239 | 65.334 | 4.448
¦- | Hg | 202 | 138 | 4.606 | 51.556 | 9.174
¦ | Pb | 208 | 1662 | 0.488 | 247.113 | 4.597
¦> | Bi | 209 | 427992 | 0.798 | 419009.917 | 1.227
¦- | Cd | 114 | 7168 | 1.182 | 225.657 | 13.932
¦> | In | 115 | 503073 | 0.722 | 491446.115 | 1.26
Concentration | Results
| Analyte | Mass | Net Intens. Mean | Conc. Mean | Conc. SD | Conc. RSD | Sample Unit | Sample Dilution Ratio | Final Sample Concentration (ppm)
¦> | Ge | 74 | 52523.438 | ppb
¦- | As | 75 | 0.02 | 0.94 | 0.01 | 1.6 | ppb | 93.9998 | 0.08836
¦- | Hg | 202 | 0 | 0.019 | 0 | 6.6 | ppb | 93.9998 | <LOQ
¦ | Pb | 208 | 0.003 | 0.067 | 0 | 1.3 | ppb | 93.9998 | <LOQ
¦> | Bi | 209 | 427992.113 | ppb
¦- | Cd | 114 | 0.014 | 0.493 | 0 | 0.5 | ppb | 93.9998 | 0.046342
¦> | In | 115 | 503073.015 | ppb
Quantitative | Analysis | - | Summary | Report
Sample | ID: | KLS-1227
Sample | Date/Time: | Thursday, June 03, 2021 23:33:29
Sample | Description: | KLS-1227
Solution | Type: | Sample
Blank | File: | C:\Path\To\Your\Data\04202021\Cal Blank.001
Number | of | Replicates: | 3
Peak | Processing | Mode: | Average
Signal | Profile | Processing | Mode: | Average
Dual | Detector | Mode: | Dual
Dead | Time | (ns): | 55
Sample | File: | C:\Path\To\Your\Data\Sample\04-20-21 CAL & SAMPLES.sam
Method | File: | C:\Path\To\Your\Data\Method\heavymetals incl is.mth
Dataset | File: | C:\Path\To\Your\Data\04202021\KLS-1227.035
Tuning | File: | C:\Path\To\Your\Data\Tuning\default.tun
Optimization | File: | C:\Path\To\Your\Data\Optimize\default.dac
Calibration | File: | C:\Path\To\Your\Data\System\04202021.cal
Calibration | Type: | External Calibration
Summary
Intensities
| Analyte | Mass | Meas. Intens. Mean | Meas. Intens. RSD | Blank Intensity | Blank Intens. RSD
¦> | Ge | 74 | 52547 | 0.383 | 47281.288 | 0.748
¦- | As | 75 | 1243 | 3.467 | 65.334 | 4.448
¦- | Hg | 202 | 118 | 7.36 | 51.556 | 9.174
¦ | Pb | 208 | 1763 | 1.494 | 247.113 | 4.597
¦> | Bi | 209 | 423284 | 0.318 | 419009.917 | 1.227
¦- | Cd | 114 | 7904 | 1.88 | 225.657 | 13.932
¦> | In | 115 | 496336 | 1.304 | 491446.115 | 1.26
Concentration | Results
| Analyte | Mass | Net Intens. Mean | Conc. Mean | Conc. SD | Conc. RSD | Sample Unit | Sample Dilution Ratio | Final Sample Concentration (ppm)
¦> | Ge | 74 | 52546.958 | ppb
¦- | As | 75 | 0.022 | 1.051 | 0.03 | 3.3 | ppb | 94.1022 | 0.098901
¦- | Hg | 202 | 0 | 0.015 | 0 | 13.6 | ppb | 94.1022 | <LOQ
¦ | Pb | 208 | 0.004 | 0.073 | 0 | 2 | ppb | 94.1022 | <LOQ
¦> | Bi | 209 | 423284.189 | ppb
¦- | Cd | 114 | 0.015 | 0.553 | 0 | 0.6 | ppb | 94.1022 | 0.052039
¦> | In | 115 | 496336.281 | ppb
Quantitative | Analysis | - | Summary | Report
Sample | ID: | Digestion Blank 04202021 + spike
Sample | Date/Time: | Thursday, June 03, 2021 23:36:34
Sample | Description: | 100 uL of 1 pmm spike - 2 ppb
Solution | Type: | Sample
Blank | File: | C:\Path\To\Your\Data\04202021\Cal Blank.001
Number | of | Replicates: | 3
Peak | Processing | Mode: | Average
Signal | Profile | Processing | Mode: | Average
Dual | Detector | Mode: | Dual
Dead | Time | (ns): | 55
Sample | File: | C:\Path\To\Your\Data\Sample\04-20-21 CAL & SAMPLES.sam
Method | File: | C:\Path\To\Your\Data\Method\heavymetals incl is.mth
Dataset | File: | C:\Path\To\Your\Data\04202021\Digestion Blank 04202021 + spike.036
Tuning | File: | C:\Path\To\Your\Data\Tuning\default.tun
Optimization | File: | C:\Path\To\Your\Data\Optimize\default.dac
Calibration | File: | C:\Path\To\Your\Data\System\04202021.cal
Calibration | Type: | External Calibration
Summary
Intensities
| Analyte | Mass | Meas. Intens. Mean | Meas. Intens. RSD | Blank Intensity | Blank Intens. RSD
¦> | Ge | 74 | 49098 | 2.846 | 47281.288 | 0.748
¦- | As | 75 | 2597 | 2.339 | 65.334 | 4.448
¦- | Hg | 202 | 9367 | 1.904 | 51.556 | 9.174
¦ | Pb | 208 | 46325 | 1.754 | 247.113 | 4.597
¦> | Bi | 209 | 463591 | 1.373 | 419009.917 | 1.227
¦- | Cd | 114 | 29282 | 3.21 | 225.657 | 13.932
¦> | In | 115 | 522981 | 1.922 | 491446.115 | 1.26
Concentration | Results
| Analyte | Mass | Net Intens. Mean | Conc. Mean | Conc. SD | Conc. RSD | Sample Unit | Final Spike Recovery %
¦> | Ge | 74 | 49097.805 | ppb
¦- | As | 75 | 0.052 | 2.431 | 0.06 | 2.3 | ppb | 121.55
¦- | Hg | 202 | 0.02 | 1.969 | 0.06 | 3.1 | ppb | 98.45
¦ | Pb | 208 | 0.099 | 2.033 | 0.06 | 2.8 | ppb | 101.65
¦> | Bi | 209 | 463591.004 | ppb
¦- | Cd | 114 | 0.056 | 1.984 | 0.03 | 1.3 | ppb | 99.2
¦> | In | 115 | 522981.075 | ppb
Quantitative | Analysis | - | Summary | Report
Sample | ID: | KLS-1228
Sample | Date/Time: | Thursday, June 03, 2021 23:39:38
Sample | Description: | KLS-1228
Solution | Type: | Sample
Blank | File: | C:\Path\To\Your\Data\04202021\Cal Blank.001
Number | of | Replicates: | 3
Peak | Processing | Mode: | Average
Signal | Profile | Processing | Mode: | Average
Dual | Detector | Mode: | Dual
Dead | Time | (ns): | 55
Sample | File: | C:\Path\To\Your\Data\Sample\04-20-21 CAL & SAMPLES.sam
Method | File: | C:\Path\To\Your\Data\Method\heavymetals incl is.mth
Dataset | File: | C:\Path\To\Your\Data\04202021\KLS-1228.037
Tuning | File: | C:\Path\To\Your\Data\Tuning\default.tun
Optimization | File: | C:\Path\To\Your\Data\Optimize\default.dac
Calibration | File: | C:\Path\To\Your\Data\System\04202021.cal
Calibration | Type: | External Calibration
Summary
Intensities
| Analyte | Mass | Meas. Intens. Mean | Meas. Intens. RSD | Blank Intensity | Blank Intens. RSD
¦> | Ge | 74 | 51524 | 0.521 | 47281.288 | 0.748
¦- | As | 75 | 1171 | 2.571 | 65.334 | 4.448
¦- | Hg | 202 | 160 | 12.686 | 51.556 | 9.174
¦ | Pb | 208 | 2315 | 0.115 | 247.113 | 4.597
¦> | Bi | 209 | 425979 | 1.688 | 419009.917 | 1.227
¦- | Cd | 114 | 4330 | 1.39 | 225.657 | 13.932
¦> | In | 115 | 493292 | 0.622 | 491446.115 | 1.26
Concentration | Results
| Analyte | Mass | Net Intens. Mean | Conc. Mean | Conc. SD | Conc. RSD | Sample Unit | Sample Dilution Ratio | Final Sample Concentration (ppm)
¦> | Ge | 74 | 51524.217 | ppb
¦- | As | 75 | 0.021 | 1.008 | 0.03 | 3.2 | ppb | 96.2873 | 0.097058
¦- | Hg | 202 | 0 | 0.025 | 0 | 16.4 | ppb | 96.2873 | <LOQ
¦ | Pb | 208 | 0.005 | 0.099 | 0 | 1.9 | ppb | 96.2873 | <LOQ
¦> | Bi | 209 | 425979.05 | ppb
¦- | Cd | 114 | 0.008 | 0.297 | 0 | 0.8 | ppb | 96.2873 | 0.028597
¦> | In | 115 | 493291.802 | ppb
Quantitative | Analysis | - | Summary | Report
Sample | ID: | KLS-1229
Sample | Date/Time: | Thursday, June 03, 2021 23:42:44
Sample | Description: | KLS-1229
Solution | Type: | Sample
Blank | File: | C:\Path\To\Your\Data\04202021\Cal Blank.001
Number | of | Replicates: | 3
Peak | Processing | Mode: | Average
Signal | Profile | Processing | Mode: | Average
Dual | Detector | Mode: | Dual
Dead | Time | (ns): | 55
Sample | File: | C:\Path\To\Your\Data\Sample\04-20-21 CAL & SAMPLES.sam
Method | File: | C:\Path\To\Your\Data\Method\heavymetals incl is.mth
Dataset | File: | C:\Path\To\Your\Data\04202021\KLS-1229.038
Tuning | File: | C:\Path\To\Your\Data\Tuning\default.tun
Optimization | File: | C:\Path\To\Your\Data\Optimize\default.dac
Calibration | File: | C:\Path\To\Your\Data\System\04202021.cal
Calibration | Type: | External Calibration
Summary
Intensities
| Analyte | Mass | Meas. Intens. Mean | Meas. Intens. RSD | Blank Intensity | Blank Intens. RSD
¦> | Ge | 74 | 55776 | 0.471 | 47281.288 | 0.748
¦- | As | 75 | 817 | 1.172 | 65.334 | 4.448
¦- | Hg | 202 | 126 | 5.524 | 51.556 | 9.174
¦ | Pb | 208 | 1442 | 3.431 | 247.113 | 4.597
¦> | Bi | 209 | 434978 | 1.74 | 419009.917 | 1.227
¦- | Cd | 114 | 1280 | 3.445 | 225.657 | 13.932
¦> | In | 115 | 531101 | 0.324 | 491446.115 | 1.26
Concentration | Results
| Analyte | Mass | Net Intens. Mean | Conc. Mean | Conc. SD | Conc. RSD | Sample Unit | Sample Dilution Ratio | Final Sample Concentration (ppm)
¦> | Ge | 74 | 55775.648 | ppb
¦- | As | 75 | 0.013 | 0.626 | 0.01 | 0.8 | ppb | 98.0735 | 0.061394
¦- | Hg | 202 | 0 | 0.016 | 0 | 12.5 | ppb | 98.0735 | <LOQ
¦ | Pb | 208 | 0.003 | 0.056 | 0 | 5.3 | ppb | 98.0735 | <LOQ
¦> | Bi | 209 | 434978.26 | ppb
¦- | Cd | 114 | 0.002 | 0.07 | 0 | 4.6 | ppb | 98.0735 | <LOQ
¦> | In | 115 | 531101.046 | ppb
Quantitative | Analysis | - | Summary | Report
Sample | ID: | KLS-1230
Sample | Date/Time: | Thursday, June 03, 2021 23:45:50
Sample | Description: | KLS-1230
Solution | Type: | Sample
Blank | File: | C:\Path\To\Your\Data\04202021\Cal Blank.001
Number | of | Replicates: | 3
Peak | Processing | Mode: | Average
Signal | Profile | Processing | Mode: | Average
Dual | Detector | Mode: | Dual
Dead | Time | (ns): | 55
Sample | File: | C:\Path\To\Your\Data\Sample\04-20-21 CAL & SAMPLES.sam
Method | File: | C:\Path\To\Your\Data\Method\heavymetals incl is.mth
Dataset | File: | C:\Path\To\Your\Data\04202021\KLS-1230.039
Tuning | File: | C:\Path\To\Your\Data\Tuning\default.tun
Optimization | File: | C:\Path\To\Your\Data\Optimize\default.dac
Calibration | File: | C:\Path\To\Your\Data\System\04202021.cal
Calibration | Type: | External Calibration
Summary
Intensities
| Analyte | Mass | Meas. Intens. Mean | Meas. Intens. RSD | Blank Intensity | Blank Intens. RSD
¦> | Ge | 74 | 51826 | 1.401 | 47281.288 | 0.748
¦- | As | 75 | 2438 | 0.699 | 65.334 | 4.448
¦- | Hg | 202 | 224 | 5.956 | 51.556 | 9.174
¦ | Pb | 208 | 846 | 0.615 | 247.113 | 4.597
¦> | Bi | 209 | 424315 | 1.037 | 419009.917 | 1.227
¦- | Cd | 114 | 6474 | 0.403 | 225.657 | 13.932
¦> | In | 115 | 488577 | 0.933 | 491446.115 | 1.26
Concentration | Results
| Analyte | Mass | Net Intens. Mean | Conc. Mean | Conc. SD | Conc. RSD | Sample Unit | Sample Dilution Ratio | Final Sample Concentration (ppm)
¦> | Ge | 74 | 51825.713 | ppb
¦- | As | 75 | 0.046 | 2.155 | 0.03 | 1.3 | ppb | 90.0507 | 0.19406
¦- | Hg | 202 | 0 | 0.04 | 0 | 6.8 | ppb | 90.0507 | <LOQ
¦ | Pb | 208 | 0.001 | 0.029 | 0 | 1.5 | ppb | 90.0507 | <LOQ
¦> | Bi | 209 | 424315.458 | ppb
¦- | Cd | 114 | 0.013 | 0.457 | 0 | 0.8 | ppb | 90.0507 | 0.041153
¦> | In | 115 | 488577.14 | ppb
Quantitative | Analysis | - | Summary | Report
Sample | ID: | KLS-1231
Sample | Date/Time: | Thursday, June 03, 2021 23:48:56
Sample | Description: | KLS-1231
Solution | Type: | Sample
Blank | File: | C:\Path\To\Your\Data\04202021\Cal Blank.001
Number | of | Replicates: | 3
Peak | Processing | Mode: | Average
Signal | Profile | Processing | Mode: | Average
Dual | Detector | Mode: | Dual
Dead | Time | (ns): | 55
Sample | File: | C:\Path\To\Your\Data\Sample\04-20-21 CAL & SAMPLES.sam
Method | File: | C:\Path\To\Your\Data\Method\heavymetals incl is.mth
Dataset | File: | C:\Path\To\Your\Data\04202021\KLS-1231.040
Tuning | File: | C:\Path\To\Your\Data\Tuning\default.tun
Optimization | File: | C:\Path\To\Your\Data\Optimize\default.dac
Calibration | File: | C:\Path\To\Your\Data\System\04202021.cal
Calibration | Type: | External Calibration
Summary
Intensities
| Analyte | Mass | Meas. Intens. Mean | Meas. Intens. RSD | Blank Intensity | Blank Intens. RSD
¦> | Ge | 74 | 51564 | 0.367 | 47281.288 | 0.748
¦- | As | 75 | 1556 | 1.808 | 65.334 | 4.448
¦- | Hg | 202 | 224 | 5.455 | 51.556 | 9.174
¦ | Pb | 208 | 754 | 1.877 | 247.113 | 4.597
¦> | Bi | 209 | 426689 | 0.2 | 419009.917 | 1.227
¦- | Cd | 114 | 2849 | 1.305 | 225.657 | 13.932
¦> | In | 115 | 496982 | 0.925 | 491446.115 | 1.26
Concentration | Results
| Analyte | Mass | Net Intens. Mean | Conc. Mean | Conc. SD | Conc. RSD | Sample Unit | Sample Dilution Ratio | Final Sample Concentration (ppm)
¦> | Ge | 74 | 51564.101 | ppb
¦- | As | 75 | 0.029 | 1.358 | 0.03 | 2.2 | ppb | 94.3934 | 0.12819
¦- | Hg | 202 | 0 | 0.039 | 0 | 7.3 | ppb | 94.3934 | <LOQ
¦ | Pb | 208 | 0.001 | 0.024 | 0 | 2.6 | ppb | 94.3934 | <LOQ
¦> | Bi | 209 | 426689.488 | ppb
¦- | Cd | 114 | 0.005 | 0.188 | 0 | 2.3 | ppb | 94.3934 | <LOQ
¦> | In | 115 | 496981.99 | ppb
Quantitative | Analysis | - | Summary | Report
Sample | ID: | Digestion Blank 04202021 3
Sample | Date/Time: | Thursday, June 03, 2021 23:52:01
Sample | Description: | Digestion Blank 04202021 3
Solution | Type: | Sample
Blank | File: | C:\Path\To\Your\Data\04202021\Cal Blank.001
Number | of | Replicates: | 3
Peak | Processing | Mode: | Average
Signal | Profile | Processing | Mode: | Average
Dual | Detector | Mode: | Dual
Dead | Time | (ns): | 55
Sample | File: | C:\Path\To\Your\Data\Sample\04-20-21 CAL & SAMPLES.sam
Method | File: | C:\Path\To\Your\Data\Method\heavymetals incl is.mth
Dataset | File: | C:\Path\To\Your\Data\04202021\Digestion Blank 04202021 3.041
Tuning | File: | C:\Path\To\Your\Data\Tuning\default.tun
Optimization | File: | C:\Path\To\Your\Data\Optimize\default.dac
Calibration | File: | C:\Path\To\Your\Data\System\04202021.cal
Calibration | Type: | External Calibration
Summary
Intensities
| Analyte | Mass | Meas. Intens. Mean | Meas. Intens. RSD | Blank Intensity | Blank Intens. RSD
¦> | Ge | 74 | 48038 | 2.045 | 47281.288 | 0.748
¦- | As | 75 | 448 | 6.189 | 65.334 | 4.448
¦- | Hg | 202 | 51 | 2.279 | 51.556 | 9.174
¦ | Pb | 208 | 165 | 5.839 | 247.113 | 4.597
¦> | Bi | 209 | 458152 | 0.487 | 419009.917 | 1.227
¦- | Cd | 114 | 170 | 8.334 | 225.657 | 13.932
¦> | In | 115 | 505174 | 2.017 | 491446.115 | 1.26
Concentration | Results
| Analyte | Mass | Net Intens. Mean | Conc. Mean | Conc. SD | Conc. RSD | Sample Unit
¦> | Ge | 74 | 48038.436 | ppb
¦- | As | 75 | 0.008 | 0.374 | 0.02 | 5.1 | ppb
¦- | Hg | 202 | 0 | -0.001 | 0 | 22.4 | ppb
¦ | Pb | 208 | 0 | -0.005 | 0 | 9.4 | ppb
¦> | Bi | 209 | 458151.979 | ppb
¦- | Cd | 114 | 0 | -0.004 | 0 | 27.2 | ppb
¦> | In | 115 | 505174.101 | ppb
Quantitative | Analysis | - | Summary | Report
Sample | ID: | KLS-1222 RR
Sample | Date/Time: | Thursday, June 03, 2021 23:55:04
Sample | Description: | KLS-1222 RR
Solution | Type: | Sample
Blank | File: | C:\Path\To\Your\Data\04202021\Cal Blank.001
Number | of | Replicates: | 3
Peak | Processing | Mode: | Average
Signal | Profile | Processing | Mode: | Average
Dual | Detector | Mode: | Dual
Dead | Time | (ns): | 55
Sample | File: | C:\Path\To\Your\Data\Sample\04-20-21 CAL & SAMPLES.sam
Method | File: | C:\Path\To\Your\Data\Method\heavymetals incl is.mth
Dataset | File: | C:\Path\To\Your\Data\04202021\KLS-1222 RR.042
Tuning | File: | C:\Path\To\Your\Data\Tuning\default.tun
Optimization | File: | C:\Path\To\Your\Data\Optimize\default.dac
Calibration | File: | C:\Path\To\Your\Data\System\04202021.cal
Calibration | Type: | External Calibration
Summary
Intensities
| Analyte | Mass | Meas. Intens. Mean | Meas. Intens. RSD | Blank Intensity | Blank Intens. RSD
¦> | Ge | 74 | 50227 | 0.481 | 47281.288 | 0.748
¦- | As | 75 | 3323 | 4.291 | 65.334 | 4.448
¦- | Hg | 202 | 79 | 2.223 | 51.556 | 9.174
¦ | Pb | 208 | 5848 | 1.346 | 247.113 | 4.597
¦> | Bi | 209 | 436223 | 3.237 | 419009.917 | 1.227
¦- | Cd | 114 | 4289 | 5.062 | 225.657 | 13.932
¦> | In | 115 | 476451 | 0.094 | 491446.115 | 1.26
Concentration | Results
| Analyte | Mass | Net Intens. Mean | Conc. Mean | Conc. SD | Conc. RSD | Sample Unit | Sample Dilution Ratio | Final Sample Concentration (ppm)
¦> | Ge | 74 | 50226.572 | ppb
¦- | As | 75 | 0.065 | 3.057 | 0.14 | 4.5 | ppb | 102.5 | 0.31334
¦- | Hg | 202 | 0 | 0.006 | 0 | 11.4 | ppb | 102.5 | <LOQ
¦ | Pb | 208 | 0.013 | 0.262 | 0.01 | 2 | ppb | 102.5 | 0.026855
¦> | Bi | 209 | 436222.897 | ppb
¦- | Cd | 114 | 0.009 | 0.305 | 0.02 | 5.4 | ppb | 102.5 | 0.031262
¦> | In | 115 | 476450.611 | ppb
Quantitative | Analysis | - | Summary | Report
Sample | ID: | QC Std 3
Sample | Date/Time: | Friday, June 04, 2021 00:01:43
Sample | Description:
Solution | Type: | QC Std
Blank | File: | C:\Path\To\Your\Data\04202021\Cal Blank.001
Number | of | Replicates: | 3
Peak | Processing | Mode: | Average
Signal | Profile | Processing | Mode: | Average
Dual | Detector | Mode: | Dual
Dead | Time | (ns): | 55
Sample | File: | C:\Path\To\Your\Data\Sample\04-20-21 CAL & SAMPLES.sam
Method | File: | C:\Path\To\Your\Data\Method\heavymetals incl is.mth
Dataset | File: | C:\Path\To\Your\Data\04202021\QC Std 3.043
Tuning | File: | C:\Path\To\Your\Data\Tuning\default.tun
Optimization | File: | C:\Path\To\Your\Data\Optimize\default.dac
Calibration | File: | C:\Path\To\Your\Data\System\04202021.cal
Calibration | Type: | External Calibration
Summary
Intensities
| Analyte | Mass | Meas. Intens. Mean | Meas. Intens. RSD | Blank Intensity | Blank Intens. RSD
¦> | Ge | 74 | 43592 | 2.414 | 47281.288 | 0.748
¦- | As | 75 | 2359 | 2.787 | 65.334 | 4.448
¦- | Hg | 202 | 10458 | 1.26 | 51.556 | 9.174
¦ | Pb | 208 | 51531 | 1.127 | 247.113 | 4.597
¦> | Bi | 209 | 464566 | 1.364 | 419009.917 | 1.227
¦- | Cd | 114 | 30278 | 0.155 | 225.657 | 13.932
¦> | In | 115 | 479083 | 0.761 | 491446.115 | 1.26
Concentration | Results
| Analyte | Mass | Net Intens. Mean | Conc. Mean | Conc. SD | Conc. RSD | Sample Unit | CCV Recovery (%)
¦> | Ge | 74 | 43591.677 | ppb
¦- | As | 75 | 0.053 | 2.489 | 0.08 | 3.3 | ppb | 99.56
¦- | Hg | 202 | 0.022 | 2.195 | 0.05 | 2.5 | ppb | 87.8
¦ | Pb | 208 | 0.11 | 2.258 | 0.04 | 1.8 | ppb | 90.32
¦> | Bi | 209 | 464565.568 | ppb
¦- | Cd | 114 | 0.063 | 2.242 | 0.02 | 0.9 | ppb | 89.68
¦> | In | 115 | 479082.928 | ppb
Quantitative | Analysis | - | Summary | Report
Sample | ID: | Digestion Blank 04202021 3
Sample | Date/Time: | Friday, June 04, 2021 00:13:08
Sample | Description: | Digestion Blank 04202021 3 RR
Solution | Type: | Sample
Blank | File: | C:\Path\To\Your\Data\04202021\Cal Blank.001
Number | of | Replicates: | 3
Peak | Processing | Mode: | Average
Signal | Profile | Processing | Mode: | Average
Dual | Detector | Mode: | Dual
Dead | Time | (ns): | 55
Sample | File: | C:\Path\To\Your\Data\Sample\04-20-21 CAL & SAMPLES.sam
Method | File: | C:\Path\To\Your\Data\Method\heavymetals incl is.mth
Dataset | File: | C:\Path\To\Your\Data\04202021\Digestion Blank 04202021 3 .044
Tuning | File: | C:\Path\To\Your\Data\Tuning\default.tun
Optimization | File: | C:\Path\To\Your\Data\Optimize\default.dac
Calibration | File: | C:\Path\To\Your\Data\System\04202021.cal
Calibration | Type: | External Calibration
Summary
Intensities
| Analyte | Mass | Meas. Intens. Mean | Meas. Intens. RSD | Blank Intensity | Blank Intens. RSD
¦> | Ge | 74 | 43716 | 0.598 | 47281.288 | 0.748
¦- | As | 75 | 404 | 7.214 | 65.334 | 4.448
¦- | Hg | 202 | 60 | 13.906 | 51.556 | 9.174
¦ | Pb | 208 | 158 | 3.106 | 247.113 | 4.597
¦> | Bi | 209 | 462975 | 0.84 | 419009.917 | 1.227
¦- | Cd | 114 | 175 | 19.443 | 225.657 | 13.932
¦> | In | 115 | 474232 | 1.289 | 491446.115 | 1.26
Concentration | Results
| Analyte | Mass | Net Intens. Mean | Conc. Mean | Conc. SD | Conc. RSD | Sample Unit
¦> | Ge | 74 | 43716.287 | ppb
¦- | As | 75 | 0.008 | 0.371 | 0.03 | 9.2 | ppb
¦- | Hg | 202 | 0 | 0.001 | 0 | 329 | ppb
¦ | Pb | 208 | 0 | -0.005 | 0 | 5.3 | ppb
¦> | Bi | 209 | 462975.256 | ppb
¦- | Cd | 114 | 0 | -0.003 | 0 | 77.6 | ppb
¦> | In | 115 | 474231.881 | ppb
Quantitative | Analysis | - | Summary | Report
Sample | ID: | KLS-1224 + POST MIC SPIKE
Sample | Date/Time: | Friday, June 04, 2021 00:19:14
Sample | Description: | KLS-1224 + POST MIC SPIKE
Solution | Type: | Sample
Blank | File: | C:\Path\To\Your\Data\04202021\Cal Blank.001
Number | of | Replicates: | 3
Peak | Processing | Mode: | Average
Signal | Profile | Processing | Mode: | Average
Dual | Detector | Mode: | Dual
Dead | Time | (ns): | 55
Sample | File: | C:\Path\To\Your\Data\Sample\04-20-21 CAL & SAMPLES.sam
Method | File: | C:\Path\To\Your\Data\Method\heavymetals incl is.mth
Dataset | File: | C:\Path\To\Your\Data\04202021\KLS-1224 + POST MIC SPIKE.046
Tuning | File: | C:\Path\To\Your\Data\Tuning\default.tun
Optimization | File: | C:\Path\To\Your\Data\Optimize\default.dac
Calibration | File: | C:\Path\To\Your\Data\System\04202021.cal
Calibration | Type: | External Calibration
Summary
Intensities
| Analyte | Mass | Meas. Intens. Mean | Meas. Intens. RSD | Blank Intensity | Blank Intens. RSD
¦> | Ge | 74 | 56663 | 0.38 | 47281.288 | 0.748
¦- | As | 75 | 3366 | 0.971 | 65.334 | 4.448
¦- | Hg | 202 | 9266 | 0.15 | 51.556 | 9.174
¦ | Pb | 208 | 53077 | 0.521 | 247.113 | 4.597
¦> | Bi | 209 | 497395 | 1.551 | 419009.917 | 1.227
¦- | Cd | 114 | 33545 | 1.331 | 225.657 | 13.932
¦> | In | 115 | 603947 | 0.932 | 491446.115 | 1.26
Concentration | Results
| Analyte | Mass | Net Intens. Mean | Conc. Mean | Conc. SD | Conc. RSD | Sample Unit | Sample Dilution Ratio | Final Sample Concentration (ppm) | Final Spike Recovery %
¦> | Ge | 74 | 56662.709 | ppb
¦- | As | 75 | 0.058 | 2.737 | 0.02 | 0.6 | ppb | 387.814 | 1.0614 | 131.6
¦- | Hg | 202 | 0.019 | 1.814 | 0.03 | 1.4 | ppb | 387.814 | 0.70349 | 90.5
¦ | Pb | 208 | 0.106 | 2.172 | 0.03 | 1.2 | ppb | 387.814 | 0.84233 | 101.75
¦> | Bi | 209 | 497395.235 | ppb
¦- | Cd | 114 | 0.055 | 1.968 | 0.02 | 0.9 | ppb | 387.814 | 0.76322 | 98.65
¦> | In | 115 | 603947.24 | ppb

## Log
Vessel | Sample ID | Sample Mass (g) | Sample Dilution | Total Mass post microwave | Total digestate mass g | Spike Amnt | HNO3/H2O | ACID LOT # | Water Lot# | Microwave | Start Time | End Time
1 | Digestion Blank 06/03/2021 | N/A | #VALUE! |  |  |  | 5/5 | HNO3-60332 | ASTM WATER - 4009 K51 | Multiwave5000 | 2021-06-03 19:00 | 2021-06-03 19:00
2 | KLS-1224 | 0.1312 | 387.814 |  |  |  | 5/5 | HNO3-60332 | ASTM WATER - 4009 K51 | Multiwave5000 | 2021-06-03 19:00 | 2021-06-03 19:00
3 | KLS-1224 + SPIKE | 0.1804 | 282.3193 |  |  | 100 uL of 1 ppm spike - 2 ppb | 5/5 | HNO3-60332 | ASTM WATER - 4009 K51 | Multiwave5000 | 2021-06-03 19:00 | 2021-06-03 19:00
4 | KLS-1225 | 0.145 | 351 |  |  |  | 5/5 | HNO3-60332 | ASTM WATER - 4009 K51 | Multiwave5000 | 2021-06-03 19:00 | 2021-06-03 19:00
5 | KLS-1226 | 0.5457 | 93.9998 |  |  |  | 5/5 | HNO3-60332 | ASTM WATER - 4009 K51 | Multiwave5000 | 2021-06-03 19:00 | 2021-06-03 19:00
6 | KLS-1227 | 0.5451 | 94.1022 |  |  |  | 5/5 | HNO3-60332 | ASTM WATER - 4009 K51 | Multiwave5000 | 2021-06-03 19:00 | 2021-06-03 19:00
7 | Digestion Blank 06/03/2021 + spike | N/A | #VALUE! |  |  | 100 uL of 1 ppm spike - 2 ppb | 5/5 | HNO3-60332 | ASTM WATER - 4009 K51 | Multiwave5000 | 2021-06-03 19:00 | 2021-06-03 19:00
8 | KLS-1228 | 0.5326 | 96.2873 |  |  |  | 5/5 | HNO3-60332 | ASTM WATER - 4009 K51 | Multiwave5000 | 2021-06-03 19:00 | 2021-06-03 19:00
9 | KLS-1229 | 0.5228 | 98.0735 |  |  |  | 5/5 | HNO3-60332 | ASTM WATER - 4009 K51 | Multiwave5000 | 2021-06-03 19:00 | 2021-06-03 19:00
10 | KLS-1230 | 0.5699 | 90.0507 |  |  |  | 5/5 | HNO3-60332 | ASTM WATER - 4009 K51 | Multiwave5000 | 2021-06-03 19:00 | 2021-06-03 19:00
11 | KLS-1231 | 0.5434 | 94.3934 |  |  |  | 5/5 | HNO3-60332 | ASTM WATER - 4009 K51 | Multiwave5000 | 2021-06-03 19:00 | 2021-06-03 19:00
12 | Digestion Blank 06/03/2021 3 | N/A | #VALUE! |  |  |  | 5/5 | HNO3-60332 | ASTM WATER - 4009 K51 | Multiwave5000 | 2021-06-03 19:00 | 2021-06-03 19:00